## Slide 1: Status and perspectives of solid state photon detectors
Sergey Vinogradov
P.N. Lebedev Physical Institute of the Russian Academy of Sciences, Moscow, Russia
Division of Solid State Physics
    National Research Nuclear University «MEPhI», Moscow, Russia
    Laboratory of Silicon Photomultipliers

## Slide 2: Scope & outline
  Photon detection and Solid State Photon Detectors (SSPD)
  Cherenkov light detection challenges
  Silicon Photomultiplier (SiPM)
    Concept
    Physics
    Designs
  SSPD / SiPM performance overview
  Photon number / energy resolution
  Time resolution
  Spatial resolution
    Position sensitive SiPM
    Digital SiPM / SPAD Arrays
  Perspectives of solid state photon detectors

## Slide 3: Cherenkov light detection challenge
Cherenkov light pulse detection objectives:
  Number of photons in the pulse
  Time of arrival (TOF) of the pulse
  Spatial distribution of the light
Cherenkov detection demands to photon detector:
  Photon number resolution – starting from single photons
  Time resolution – pico-…nanosecond range
  Spatial resolution – (micro-?)… mm range
Ideally, we prefer the highest precision
in measuring #photons in time and space
In fact, Cherenkov detection is
an ultimate challenge for photon detectors
Photons
Images
E. Fossum, Sensors, 2016

## Slide 4: Why Solid State Photon Detectors?
Advantages of solid state (semiconductor) photodetectors
  High photon detection efficiency (PDE) due to internal photoemission
    For good photon number resolution
  Fast timing due to µm-scale thickness of sensitive layers
    For good time resolution
  Scalable active areas in µm – mm range due to solid state technology
    for good spatial resolution
  Solid state technology
Trade-offs
  Fast timing => high electronic noise
  ~ 10K electrons @ 1 GHz BW
  => multiplication with gain > 10K
  for single photon detection => APD
  High gain APD =>
  high excess noise of avalanche
  no photon number resolution
W. Farr, SPIE LEOS 2009
New kind of APD with high gain, low noise, fast timing is required:
Solid State / Silicon Photomultiplier (SSPM / SiPM)

## Slide 5: SiPM: concept, physics, designs. developments
SiPM/SSPM: concept, physics, designs. developments

## Slide 6: SiPM concept: Basic studies at Lebedev Physical Institute (1970s-1990s)
Avalanche with Negative Feedback (ANF) concept:
  Strong positive feedback: avalanche development as fast as possible
  → Geiger mode avalanche breakdown with high overvoltage
  Strong negative feedback: avalanche quenching as fast as possible
  → Passive quenching by built-in elements (ANF or NAF APD)
  Capacitive (feedback ~ avalanche charge)
  Resistive (feedback ~ avalanche current)
V. Shubin, D. Shushakov, Encyclopedia of Optical Engineering 2003; DOI: 10.1081/E-EOE 120009727
ANF
=>
High gain
+
Low noise
+
Fast time

## Slide 7: SiPM physics: model of single-electron ANF process
Distribution of single electron Gain Excess Noise Factor of Gain
Gain vs Overvoltage ANF process time
D. Shushakov, V. Shubin, New Solid State Photomultiplier, SPIE (1995)

## Slide 8: SiPM physics: experiment on single-electron ANF process
D. Shushakov, V. Shubin, New Solid State Photomultiplier, SPIE (1995)
A few photon pulse detection by charge sensitive preamp
Experimental histograms of single photoelectron Gain
Pedestal
1 phe
Simulated distribution of single electron Gain

## Slide 9: SiPM designs: Basic developments (1988-1998)
MOS APD (1970s-1980s)
Planar MRS APD (1988)
Needle MRS APD (1989)
MRS APD with trenches (1998)
Z. Sadygov, NDIP 2005, 2008; ICASiPM 2018
Key developers:
Z. Sadygov (INR, JINR)
V. Golovin (MELZ, CPTA)

## Slide 10: SiPM designs: Advanced developments (2000s)
MOS APD (1970s-1980s)
SiPM, B. Dolgoshein et al., MEPhI / Pulsar, 2000-2003
SSPM, V. Golovin et al., CPTA / Photonique, NDIP 2008
Micropixel APD, Z. Sadygov et al., JINR / Zecotek, NDIP 2005

## Slide 11: Modern SiPMs
SiPM/SSPM – a new generation of photon-number-resolving SSPDs
Analog SiPM/SSPM – an array of APD cells with individual quenching resistors and common readout independently operated in a limited Geiger mode
Alternative SiPM designs: Micropixel APDs, SiPMs with bulk-quenching…
Convergence of SiPM technology and SPAD arrays with active quenching of CMOS technology: digital SiPM, multi-digital SiPM, 3D SiPM…
KETEK SiPM – custom technology STMicroelectronics SiPM – CMOS technology

## Slide 12: SiPM: performance
SiPM/SSPM: performance in photon number resolution

## Slide 13: SiPM: photon number resolution
B. Kardinal et al., Nat. Photonics, 2008
R. Mirzoyan et al., NDIP, 2008
A. Barlow and J. Schilz, SiPM matching event, CERN, 2011
APD (self-differencing mode)
SiPM (MEPhI/Pulsar)
SiPM (Excelitas by MPI/MEPhI license)
PMT (Hamamatsu H11284-100UV)
T. Cogami et al, Water Cherenkov detector, NIMA 2016
S. Vinogradov, SPIE 2017
SiPM (Hamamatsu MPPC)
Charge integration
Charge histogram
APD QE ~ 80% SiPM PDE ~ 30%
PMT QE ~ 30%
SiPM (Hamamatsu MPPC)

## Slide 14: Excess noise factor (ENF) as a measure of noisiness of multipliers and amplifiers
ENF in multipliers (APD, PMT, SiPM):
ENF in amplifiers:
Noise factor (F) measures degradation of the signal-to-noise ratio (SNR), caused by components in a signal chain.
High ENF
Low ENF
1 electron → M electrons

## Slide 15: Resolution of photon detection: Excess Noise Factor approach[1]
  Output RES defined by input RES degraded due to a photodetection processes with specific ENFs
Input photon distribution Photodetection processes Output charge distribution
RES(Xin)
RES(Yout)
Process i
ENFi
Process j
ENFj
[1] S. Vinogradov et al., “Efficiency of Solid State Photomultipliers in Photon Number Resolution,” IEEE Trans. Nucl. Sci., 2011.
[2] S. Vinogradov, “Analytical models of probability distribution and excess noise factor of solid state photomultiplier signals with crosstalk,” NIMA 2012.

## Slide 16: Key SiPM parameters in ENF metrics

## Slide 17: Resolution, total ENF, and Detective Quantum Efficiency (DQE)
DCR
Npix, Trecov
Max DQE

## Slide 18: Performance of photodetectors in DQE
Electronic & Dark noise
Nonlinearity
Saturation
Optimal detection
Detective Range
Optimal DQE

## Slide 19: The highest max DQE ~ 55% (2011) PDE ~ 58%, Crosstalk ~ 4%
E. Popova et al., NDIP 2011

## Slide 20: Highest DQE of market SiPMs (2014)
Short light pulse detection (Tpulse < Tdead), binomial nonlinearity
S. Vinogradov, 2nd Advanced SiPM Workshop, CTA, Geneva, 2014
Optimal DQE
(2014)

## Slide 21: DQE trade-offs: PDE vs Crosstalk vs Dynamic Range
Short light pulse detection (Tpulse < Tdead), binomial nonlinearity
S. Vinogradov, 2nd Advanced SiPM Workshop, CTA, Geneva, 2014
100um: high PDE=37% but high Pct=35%
10um: high Dynamic Range but low PDE=10% & low Cpix
low Gain → electronic noise
(2014)

## Slide 22: DQE trade-off: PDE vs Crosstalk vs Dynamic Range
General: High PDE = QE · Pav_trig · Fill Factor (with QE, Pav_trig → 90%)
Direction 1: Large pixels (50…75 um) with high fill factor and max suppression of crosstalk (metal-filled trenches and more). Max priority – high DQE for low photons. Low Dynamic range as is.
Direction 2: Small pixels (5…15 um) with low crosstalk due to low gain (low capacitance) and fast recovery (low capacitance). Max priority – high Dynamic range. DQE as high as possible (no trenches, narrow pixel borders for higher fill factor)
DCR
Npix, Trecov
Crosstalk
PDE

## Slide 23: High PDE: QE · Pav_trig · Fill Factor

## Slide 24: QE in UV: MPPC for Cherenkov and VUV detection
S14520 Series (50 um) – developed for Cherenkov Telescope Array
S13370 series – developed for VUV
Cryogenically reliable
Low optical crosstalk due to trenches
Fast recovery (stable Rq vs. temperature)

## Slide 25: High PDE: QE · Pav_trig · Fill Factor
=> Pav_trig ~ 95%

## Slide 26: High PDE: QE · Pav_trig · Fill Factor

## Slide 27: Large pixels with trenches: low crosstalk and high fill factor
F.Wiest NDIP2014

## Slide 28: Small pixels with low gain: low crosstalk and high dynamic range
F.Acerbi, TREDI 2017

## Slide 29: Highest DQE of modern SiPMs (2018)

## Slide 30: SiPM: performance
SiPM/SSPM: performance in time resolution

## Slide 31: SiPM time resolution
Delta-pulses - Single Photon Time Resolution (SPTR) (+ +)
  SiPM characterization; Time-Correlated Single-Photon Spectroscopy
Gaussian pulses of variable width (laser/LED) (+)
  LIDAR
Scintillation pulses – Coincidence Time Resolution (+ + +)
  TOF PET and more…
Arbitrary light waveforms – almost out of SiPM studies (-)
  Cherenkov Fiber Beam Loss Monitoring
Leading edge
or
Constant fraction discriminator
Single-electron response (SER) shape
S. Vinogradov, NDIP 2014, NIMA 2015

## Slide 32: Main trade-off: fast timing vs large area
Larger area => higher noise + slower response => lower TR
  DCR noise ~ area
  Electronic noise ~ capacitance ~ area
  Signal rising slope ~ SER rise time ~ capacitance
F.Acerbi, FAST 2014, NIMA 2015

## Slide 33: Compensation of large capacitance by transformer

## Slide 34: SPTR and CTR of modern SiPMs to be reconsidered

## Slide 35: Strip SiPM with dual differential readout
K. Doroud et al, NIMA 2017

## Slide 36: Performance of Strip SiPM
SPTR ~ 70 ps due to dual differential readout
  Typical SPTR of 3x3mm2 SiPM ~ 200 ps
Position resolution by time delay encoding (50 ps) ~ 750 um
K. Doroud et al, NIMA 2017

## Slide 37: SiPM: performance
SiPM/SSPM: approaches to high spatial resolution
    Position Sensitive SiPMs
    SPAD arrays
    Quantum Image Sensors

## Slide 38: Position-Sensitive SiPM
D. Han, ICASiPM 2018

## Slide 39: 3x3mm2 PS SiPM of Novel Device Lab (Beijing)
High position resolution (26…40 um)
High photon number resolution
D. Han, ICASiPM 2018

## Slide 40: SPAD arrays trade-off: TDC multiplexing

## Slide 41: SPAD arrays trade-off: PDE vs functionalities
J. F. Pratte et al., FAST 2014

## Slide 42: Digital SiPM Philips digital photon counter
Fill Factor ~ 50%
PDE ~ 30%
SPRT~ 85 ps for “pixel” of 3200 cells (3.2 x 3.9 mm2)
SPTR ~ 100 ps for SiPM of 4 “pixels”
Readout time = 327 µs
T. Frach et al., IEEE NSS/MIC 2009

## Slide 43: SPAD arrays: Multi-Digital SiPM and more… Aqua Lab EPFL / TU Delft (Edoardo Charbon)
E. Charbon, FAST 2018

## Slide 44: Performance of SPADs
E. Charbon, FAST 2018

## Slide 45: 3D SiPM (SPAD Array), Sherbrook University
J. F. Pratte et al., FAST 2014
F. Nolet et al., NDIP 2017
SPTR of single SPAD + TDC = 17.5 ps FWHM

## Slide 46: Quantum Image Sensor: ultimate CMOS imager Dartmouth College (Eric Fossum)
Size < 500 nm
Capacitance < 1 fF
Gain ~ 1mV/e
Noise < 0.15 e RMS
J. Ma et al., Photon-number-resolving megapixel image sensor at room temperature without avalanche gain, Optica 2017

## Slide 47: Summary
SiPM technology: breakthrough in photon detection
  Superior photon number resolution and time resolution
  Highly competitive in short (ps – ns range) pulse detection
  CMOS-compatible / mass production / low cost / reliable technology
  Position-sensitive SiPMs are promising as inexpensive imagers
  Fast progress in:
    PDE
    Dynamic range
    SPTR
    Non-silicon SSPM
      NIR
      UV
      RadHard
  Convergence of
    Analog SiPMs
    Digital SiPMs
    SPAD arrays
    3D SiPMs

## Slide 48: Acknowledgements
This work was supported in part by EC under
Grant Agreement FP7 329100 “SiPM in-depth”
FAST (Fast Advanced Scintillation Timing) COST TD1401 action
“SENSE – a roadmap for the ideal low light level sensor development” Horizon 2020 FET-Open CSA project

## Slide 49: The end
Thank you for your attention!
Questions?
Objections?
Opinions?
…
vin@lebedev.ru

## Slide 51: SPAD arrays: Multi-Digital SiPM and more…
E. Charbon, FAST 2018